## Input
Signal | Direction | Type | Quantity | Properties | Status | Description
MP | In | TTL | 1 | Pulse (typical <=50 Hz, <=10ms ) |  | Marcopulse, in future 10ms
DT2 | In | TTL | 1 | Frequeny <8MHz |  | Trafo Current
DT2Range | In | TTL | 3 | Range ID (3Bit) |  | Measurement range
DT3 | In | TTL | 1 | Frequeny <8MHz |  | Trafo Current
DT3Range | In | TTL | 3 | Range ID (3Bit) |  | Measurement range
Detector | In | TTL | 1 | Pulse <4kHz |  | Count rate phsiacl trigger
Rutherford | In | TTL | 2 | Pulse <5kHz |  | Count rate
TWSegmentID | In | TTL | 3 | bit <200Hz, Width |  | Coding target wheel segments
TWSync | In | TTL | 1 | Level | High:Synch; Low:Asynch | Status
Event1 | In | TTL | 1 | Pulse <50ns |  | Nuclear reaction event in energy window 1
Event2 | In | TTL | 1 | Pulse <50ns |  | Nuclear reaction event in energy window 2
PyrometerLimit | In | TTL | 1 | Level | High: T>Limit | Temperature limit exceeded
PyrometerTemp | In | TTL | 1 | Frequeny <10kHz | Temp: U/F Wandler | Temperature
FSV (X8) | In | TTL | 1 | Level | High:open; Low:closed | Close "Fast shutting valve" Tasca
SSV (Unilac) | In | TTL | 1 | Level | High:open; Low:closed | Fast shutting valves Unilac
SVHV | In | TTL | 1 | Level | High:open; Low:closed | Close Valve SV345
SlitPositionLeft | In | 4-20mA | 1 | Position |  | Absolute position of left slit
SlitPositionRight | In | 4-20mA | 1 | Position |  | Absolute position of right slit
Magnet_rigid_D | In | TTL | 1 | Freq < 10KHz |  | Magnetic Rigidity_Dipol
Magnet_rigid_Q1 | In | TTL | 1 | Freq < 10KHz |  | Magnetic Rigidity_Quadrupole1
ExternalChemistry | In | Relais |  | Level | Low: IL active | on/off Chemistry-> inhibit the beam
IL_TW | In | TTL | 1 | Level | Low: IL active | Interlock signal from target wheel
IL_Magnet | In | TTL | 1 | Level | Low: IL active | Interlock signal from magnet control
IL_FSV | In | TTL | 1 | Level | Low: IL active | Interlock signal from fast shutting valves
IL_Manual | In | TTL | 1 | Level | Low: IL active | Interlock signal set by user with button
IL_TPG300 | In | Relais | 12 | Level | Low: IL active | Interlock signal from TPG300 relais
IL_MKS | In | Relais | 8 | Level | Low: IL active | Interlock signal from MKS647 relais
IL_Rutherford | In | TTL | 1 | Level | Low: IL active | Interlock signal from Rutherford detector
IL_VetoDetector | In | Relais | 4 | Level | Low: IL active | Interlock signal from Veto detector
IL_Hydrogen | In | Relais | 4 | Level | Low: IL active | Interlock signal from Hydrogen monitor
IL_External | In | TTL | 8 | Level | Low: IL active | Interlock signal set by external systems
Number of signals: |  |  | 68

## Output
Signal | Direction | Type | Quantity | Properties | Status | Description
IL_Chopper | Out | TTL | 1 | Level | Low: IL active | Interlock signal sent to chopper to switch off the ion beam
Shutoff TASCA gas | Out | TTL | 1 | Level | Low: IL active | close Tasca gas supply when FSVs and HSs/SVs closed
SumBeamInt | Out | TTL | 1 | Level | Low: IL active | stopped e.g. chemistry when beam integral is reached
IL_External | Out | TTL | 8 | Level | Low: IL active | Interlock signal to be used by external systems
DAQ_Active | Out | TTL | 1 | Level | Low: IL active | Inticator that a measurement is active
Number of signals: |  |  | 12